# The EEZ BB3 overview
## Front panel overview

Fig. 1: Left side of the EEZ BB3 chassis front panel
	4.3” (480 x 272) color TFT touch-screen display, 16-bit (65536) colors with brightness and luminosity control: used for local system and module specific parameter display and settings.
	AC mains power switch: turns the EEZ BB3 on and off.
	Standby indicator: on when EEZ BB3 is in standby mode (power modules AC mains power is switched off).
	Protective Earth (PE) 4 mm socket.
	Master MCU BOOT0 miniature tactile switch access hole: when pressed on start up the EEZ BB3 enters a special bootloader mode for system firmware uploading.
	RJ-45 Ethernet socket (10/100 Mbit/s)
	Digital I/O terminals (two input and two output, protected)
	Micro SD card with card detection switch.
	5-pin USB 2.0 Mini AB socket
	Incremental encoder with switch
	User-defined tactile switch can be assigned to perform selected action e.g. entering inhibit mode, taking screenshot, manual triggering, etc. (see User SW).

## Module front panels overview
The EEZ BB3 can accommodate up to three modules. An example with three power modules are shown on Fig. 2.

Fig. 2: Right side of the EEZ BB3 chassis front panel
	Operation mode indicators: CC – Constant Current mode, CV – Constant Voltage mode, Rprog – Remote programming is active (DCP405 only), Sense – Remote sensing is selected (DCP405 only)
	Advanced options input terminals (DCP405 only): Remote programming, Remote sensing and PE. If PE is used for shielding remote sensing cable, connect it on one side only to avoid creation of ground loop.
	Positive output (Vout+) power connector, 4 mm, “floating” (i.e. galvanically isolated from the MCU module ground, therefore not connected to the PE potential). / Warning: The potential on this point could be up to 80 V on slot #1 when two power modules are coupled in series.
	OE (Output Enable) indicator, bi-color: Green – output is active and uncoupled, Red – output is active and coupled.
	Negative output (Vout-) power connector, 4 mm, “floating”
	Positive output (Vout+) power connector, 4 mm (DCM220 only)
	Channel indicator, bi-color (DCM220 only): Green – output is active in CV mode, Red – output is active in CC mode.
	Negative output (Vout-) power connector, 4 mm (DCM220 only). Please note that Vout- of Ch1 and Ch2 are on the same potential but still “floating” (i.e. not connected to PE potential).

## Rear panel overview

Fig. 3: The EEZ BB3 rear panel
	Low noise Ø80 mm cooling fan with speed algorithm based on multiple temperature sensors and output current readings.
	Kensington Security Slot™ (K-slot) for anti-theft cable.
	AC power inlet with two 20 x 5 mm fuses

## Display pages overview
The color display is used as a primary means of user interaction. Its content is dynamically rendered in order to enable quick and structured access to a multitude of system and module specific parameters. The touch-screen completely eliminates the need for specialized function keys, dedicated numeric keypad, etc. It takes no more than a couple of clicks to access any function, prioritizing access to those functions that are more important or used more frequently.

Fig. 4 shows home page of the EEZ BB3 loaded with three different power module (see Fig. 2) using the default view (numeric) for presenting output states, and measured and set values. This module configuration (i.e. DCP405, DCP405B and DCM220) will be used as an example throughout this manual.

Items  to  are module specific while the rest (displayed on the status bar) are used for accessing various system options.

Fig. 4: Home page overview
	Channel header: for better distinction between modules different colors are used for each header. Module name and assigned channel number are displayed together with shortcuts to channel settings (a shortcut to set or clear protection modes – If any have been tripped – and a channel “maximize” button).
	Channel main display section contains information about measured output voltage and current. Units of measurement are also displayed depending on the measured value (i.e. mV or V for voltage or μA, mA or A for current. / Note – this section looks different for 2-channel modules such as DCM220 due to more limited display area.
	Channel supplemental display section is used for set output voltage, current and power. Additionally is used for displaying measured output power together with related units of measurement (mW or W). / Note – this section looks different for 2-channel modules such as DCM220 due to more limited display area.
	Event view: View the event log. The icon color will change color if unread warning or error messages have been added to the log.
	USB activity indicator and shortcut to serial communication settings. The icon is not visible if USB/serial communication is disabled and its background becomes highlighted if USB/serial communication is established.
	LAN activity indicator and shortcut to the Ethernet communication settings. The icon is not visible if Ethernet communication is disabled and its background becomes highlighted if LAN communication is established.
	Recording / Stop: Data logger control. Use to start data logging or to stop logging before a defined end.
	Data log view: if logging is in progress it displays real time data. When logging is completed or stopped, it will display data from the last log.
	File manager: Allows access to data on a SD card such as user profiles, program lists, logged data or screenshots. Logged data and screenshots can be uploaded to EEZ Studio with one click.
	User profiles: 10 different profiles (with auto-recall function) are available for saving and recalling module parameters.
	Tracking and coupling settings: tracking allows simultaneous programming of two or more module's output channels. Coupling can be used to safely connect (under firmware control) two or more module's outputs in series, parallel, as “split rail” (i.e. series with common output) or with Vout- tied together. No additional external connection wiring is required, and in case of coupling in series the output voltage range is doubled (e.g. to 80 V). Similarly, when parallel coupling is performed the output current range is doubled (e.g. to 10 A).
	Display view selection: toggle between five different presentations of output module parameters (default is numeric view – as shown on Fig. 4).
	Display lock/unlock: when locked all local user interaction via touch-screen display will be disabled until the same option is selected again. Unlocking is not possible without entering valid system password (the system password is not defined by default).
	System settings: use it to access system-wide protection settings, display, sound, trigger, digital I/O, date & time, Encoder and user switch, AUX temperature & fan control, Ethernet, Serial/USB and channels calibration
	Power / Reset control: allows access to standby, “soft” reset (see also *RST), “hard” reset (equivalent to power up reset) and option to turn off display

## Module display views
Five different views are available to satisfy various use cases as well as user preference for numeric or graphic representation of measured output values.

Regardless of the selected display view and color theme each channel will be represented in consistent order with its assigned color. The same convention is also used when a channel is displayed over the whole screen (e.g. in the case of Protection mode settings or the channel settings page). In those cases where a channel is maximized the channel color is visible on the status bar instead of the header area.

Fig. 5: Display view icon
Use icon shown on Fig. 5 to toggle between all available display views.

Numeric / Default display view with the largest font used to show measured voltage and current. / Output values of 2-channel modules like DCM220 are displayed with a regular font size due to lack of space. / SCPI / DISPlay:VIEW 1

Vertical bar / This view is useful when the output values are changing rapidly. / Currently measured values are displayed with regular font size, hence user has to be closer to the EEZ BB3 to read the displayed values. / SCPI / DISPlay:VIEW 2

Horizontal bar / Similarly to previous view but output values are displayed as horizontal bars. / This view allows graphical presentation of the 2-channels modules, too. / SCPI / DISPlay:VIEW 3

YT view (scroll) / Measured values are displayed on the Y-axis of a 2D graph. The cursor position moves with the speed defined by YT view sampling rate (default is 100 ms). / The cursor moves cyclically – rewriting older data with new data. / SCPI / DISPlay:VIEW 4
YT view (scan) / Similar to the previous view but a moving cursor is not used, and all data displayed moves from right to left with the speed defined by YT view sampling rate (default is 100 ms). / This type of presentation is also known as roll mode. / SCPI / DISPlay:VIEW 5