# The EEZ BB3 overview
## Front panel overview

Fig. 1: Left side of the EEZ BB3 chassis front panel
	4.3” (480 x 272) color TFT touch-screen display, 16-bit (65536) colors with brightness and luminosity control: used for local system and module specific parameter displaying and settings.
	AC mains power switch: turning the EEZ BB3 on and off.
	Standby indicator is turned on when EEZ BB3 is in standby mode (modules AC mains power is switched off).
	Protective Earth (PE) 4 mm socket.
	Master MCU BOOT0 miniature tactile switch access hole: when used on start up causes starting of bootloader for system firmware uploading.
	RJ-45 Ethernet socket (10/100 Mbit/s)
	Digital I/O terminals (two input and two output, protected)
	Micro SD card with card detection switch.
	5-pin USB 2.0 Mini AB socket
	Incremental encoder with switch
	User-defined tactile switch can be assigned to perform selected action e.g. entering inhibit mode, taking screenshot, manual triggering, etc. (see User SW).

## Module front panels overview
The EEZ BB3 can be configured in different ways so that the number of each module can vary from zero to three. An example with three power modules are shown on Fig. 2.

Fig. 2: Right side of the EEZ BB3 chassis front panel
	Operation modes indicators: CC – Constant Current mode, CV – Constant Voltage mode, Rprog – Remote programming is active (DCP405 only), Sense – Remote sensing is selected (DCP405 only)
	Advanced options input terminals (DCP405 only): Remote programming, Remote sensing and PE. If PE is used for shielding remote sensing cable, connect it on one side only to avoid creation of ground loop.
	Positive output (Vout+) power connector, 4 mm, “floating” (i.e. galvanically isolated from the MCU module ground, therefore not connected to the PE potential). / Potential on this point could be up to 80 V on slot #1 when two power modules are coupled in series.
	OE (Output Enable) indicator, bi-color: Green – output is active and uncoupled, Red – output is active and coupled.
	Negative output (Vout-) power connector, 4 mm, “floating”
	Positive output (Vout+) power connector, 4 mm (DCM220 only)
	Channel indicator, bi-color (DCM220 only): Green – output is active in CV mode, Red – output is active CC mode.
	Negative output (Vout-) power connector, 4 mm (DCM220 only). Please note that Vout- of Ch1 and Ch2 are on the same potential but still “floating” (i.e. not connected to PE potential).

## Rear panel overview

Fig. 3: The EEZ BB3 rear panel
	Low noise Ø80 mm cooling fan with speed algorithm based on multiple temperature sensors and output current readings.
	Kensington Security Slot™ (K-slot) for anti-theft cable.
	AC power inlet with two 20 x 5 mm fuses

## Display pages overview
The color display is used as a primary mean or local interaction with the user. Its content is dynamically rendered in order to enable quick and structured access to multitude of system and module specific parameters. The touch-screen completely eliminated the need for specialized function keys, dedicated numeric keypad, etc. It now takes no more than a couple of clicks to get to any function, prioritizing access to the functions that are more important or used more frequently.
Fig. 4 shows home page of the EEZ BB3 loaded with three different power module (see Fig. 2) using the default view (numeric) for presenting output states, and measured and set values. This module configuration (i.e. DCP405, DCP405B and DCM220) will be used as an example throughout the text of this manual.

Items  to  are module related and the rest, displayed on the status bar, are used for accessing various system options.

Fig. 4: Home page overview
	Channel header: for better distinction between modules different colors were used for each header. Module name and assigned channel number is displayed together with shortcuts to channel settings (toggles to protection clear if any protection was tripped) and maximize channel.
	Channel main display section contains information about measured output voltage and current. Units of measurement are also displayed depending on the measured value (i.e. mV or V for voltage or μA, mA or A for current. / This section looks differently for 2-channel modules such as DCM220 due to more limited displaying area.
	Channel supplemental display section is used for set output voltage, current and power. Additionally is used for displaying measured output power together with related units of measurement (mW or W). / This section looks differently for 2-channel modules such as DCM220 due to more limited displaying area.
	Event view: use it to access event log. The icon color will change color if unread warning or error message has been added into the event log.
	USB activity indicator and shortcut to serial communication settings. The icon is not visible if USB/serial communication is disabled and its background becomes highlighted if USB/serial communication is established.
	LAN activity indicator and shortcut to the Ethernet communication settings. The icon is not visible if Ethernet communication is disabled and its background becomes highlighted if LAN communication is established.
	Tracking and coupling settings: tracking allows simultaneous programming of two or more module's output channels. Coupling can be used to connect safely (under firmware control) two or more module's outputs in series, parallel, as “split rail” (i.e. series with common output) or with Vout- tied together. No additional external connection wiring is required, and in case of coupling in series the output voltage displaying and programming values could be doubled (e.g. up to 80 V). Similarly, when parallel coupling is performed the output current displaying and programming values could be doubled (e.g. up to 10 A).
	User profiles: 10 different profiles (with auto-recall function) are available for saving and recalling module parameters.
	System information: shortcut to page that contains overview of various system information such as activity counters, fan, battery, serial number and firmware version, SD card status and installed module.
	Display view selection: toggle between five different presentation of output module parameters (default is numeric as shown on Fig. 4).
	Display lock/unlock: when locked all local user interaction via touch-screen display will be inhibited until the same option is selected again. Unlocking is not possible without entering valid system password (the system password is not defined by default).
	System settings: use it to access system-wide protection settings, display, sound, trigger, digital I/O, date & time, Encoder and user switch, AUX temperature & fan control, Ethernet, Serial/USB and channels calibration
	Power / Reset control: allows access to standby, “soft” reset (see also *RST), “hard” reset (equivalent to power up reset) and option to turn off display

## Module display views
Five different views are available to meet different user scenarios as well as user preferences for primarily numerical or graphical representation of the measured output values.

Regardless of the selected display view and color theme every channel will be represented on the same position with different color. That convention is also used when channel is displaying over the whole screen like in case of Protections and settings page. In that case channel color is applied on the status bar instead of header area.

Fig. 5: Display view icon
Use icon shown on Fig. 5 to toggle between all available display views.

Numeric / Default display view with the largest font used in presenting measured voltage and current. / Output values of 2-channel modules like DCM220 are displayed with regular font size due to lack of space. / SCPI / DISPlay:VIEW 1
Vertical bar / This view is useful when the output values are changing with greater dynamics. / Currently measured values are displayed with regular font size, hence user has to be closer to the EEZ BB3 to read the displayed values. / SCPI / DISPlay:VIEW 2
Horizontal bar / Similarly to previous view but output values are displayed as horizontal bars. / This view allows graphical presentation of the 2-channels modules, too. / SCPI / DISPlay:VIEW 3

YT view (scroll) / Measured values are displaying on Y-axis on the 2D graph on the cursor position that is moved with the speed defined by YT view sampling rate (default is 100 ms). / Cursor is moving cyclically rewriting old displayed data with the new one. / SCPI / DISPlay:VIEW 4
YT view (scan) / Similar to previous view but moving cursor is not used, and all data displayed are starts to move from right to left with the with the speed defined by YT view sampling rate (default is 100 ms). / This type of presentation is also known as roll mode. / SCPI / DISPlay:VIEW 5