# Getting started
## Powering on
Before switching on check that Safety Requirements actions are observed. Connect the power cable to the AC power connector on the rear panel and switch on AC mains power switch on the front panel. The welcome page will appear (Fig. 1) and self-test procedure is performed when firmware will try to detect and recognize all installed modules and diagnose if any error exists that could inhibit normal operation.

Fig. 1: Welcome page (light color theme)
On the first power up all channel outputs are turned off to prevent connected loads from being damaged unintentionally.
The channel output state at all subsequent power-up will depend on what is defined in the selected user profile. A special option in System settings can be set to override states written in the selected user profile.
Both default display view is numeric, and color theme that is set to dark can be changed. the home page of our 4-channel configuration example will look as shown in Fig. 2:

Fig. 2: Home page on first start
## Switching off and standby mode
The EEZ BB3 offers two option to stop its operation: turn it off using AC power switch on the front panel or by entering standby mode. In both case all installed modules are cut off from AC power. Standby mode keeps the MCU module powered that is indicated with the Standby indicator on the front panel. Entering the standby mode is shown on Fig. 3 that followed by the info message on the screen that is fading out.
Exit from the standby mode is possible by tap any part of the display and wait for a moment to power up procedure started.

Fig. 3: Entering standby mode
 SCPI
SYSTem:POWer OFF

# Basic operations
## Turn channel output on and off
Turn channel on (OE on) / Turning the channel output on is performed by taping anywhere inside / the channel main display section. / Successfully changed output state will be acknowledged by replacing the “OFF” with measured output values. / SCPI / OUTPut ON / OUTPut ON, <channel>
Turn channel off (OE off) / Turning the channel output off is performed by taping anywhere inside / the channel main display section. The content of that section will be replaced with the “OFF” sign. / SCPI / OUTPut OFF / OUTPut OFF, <channel>

## Set channel output values
Channel output values can be set regardless of the currently selected display view and channel output state as long as trigger is not active (i.e. channel’s trigger is set to Fixed).
When desired output parameter is selected, as in example below, it is possible to start changing existing value immediately by using the encoder knob on the front panel. By default, encoder switch is set to, when pressed, make next output parameter selected. That operation is indicated by highlighting parameter’s background area.

	Set output voltage / Tap inside the voltage area of / the channel supplementary display section. If its background is highlighted as on the picture, a keypad will appear as described in Data input methods. / The set voltage represents the max. output voltage when channel is in CV mode. / SCPI / VOLTage <voltage>
	Set output current / First tap inside the current area of the channel supplementary display section. Second tap is required if this option was not selected previously (as in above example). / The set current represents the max. output current when channel is in CC mode. / SCPI / CURRent <current>
	Set output power limit / Follow the above described procedure. / Note that when power limit is set, product of set voltage and current is checked, not their measured values. For example, if power limit is set to 100 W (DCP405) then set current cannot exceed 2.5 A for max. voltage (40 V). Similarly, if the output voltage cannot be set above 20 V if max. current is selected (5 A) even if output measured (or expected) current with connected load is well below set limit. / SCPI / POWer:LIMit <power>

## Data input methods
There are three ways of entering data in addition to the encoder that is already mentioned. Additionally, two of them (Step and Slider) could be used in non-confirmation (default) and confirmation mode and in combination with the encoder knob. All data input methods are accessible from the same popup menu that has three tabs on the right side.

Keypad
	This is the default data entry method, and in case that another one is currently selected it can be selected again by tap on its tab icon. / When selected, simply enter new numeric value for the parameter that is displayed in the header (Ch1 voltage in this example). / You can also select unit of measurement if more then one is available (e.g. mV and V for voltage). / This data entry method always requires confirmation.
Step
	Tap on step method icon.
	Select step value.
	Increase or decrease set value of the selected parameter in selected steps. If encoder knob is used the selected parameter value will be changed for in selected step increments (1 V in this example).
Slider
	Tap on slider method icon.
	Increase or decrease set value of the selected parameter.
Confirmation mode
	Tap to display confirmation options on the bottom. It can be used in Step and Slider mode. / When selected parameter is changed its current value will starts to blink until it is confirmed with Set option when it become effective. The new parameter value can be reset to the previous value using the Discard option.

## Accessing channel protections and settings
The EEZ BB3 firmware include comprehensive control of each power module and feature list goes well beyond basic operations of setting and displaying output parameters. All channel specific functions could be found under Protections and settings page.

The channel Protections and settings page can be reached whenever is channel in regular and maximized view.

Note that if any of channel’s protection tripped the Protection icon will lead to Clear / Disable protection page regardless is channel is in regular of maximized view.

	Access protections and settings in regular view / Select the channel whose settings you want to change and tap on its Protection icon. / A new page with channel’s protection section and channel settings will be displayed. / Availability of certain options depend of the channel mode (e.g. when channel is in tracking or coupling mode Calibration is disabled) and capability (e.g. for example the DCP405 and DCM220 modules do not have Advanced options.
	Exit Protections and settings page.
	Access protections and settings from maximized view / If channel is in regular view, tap on the channel’s Maximize icon.
	Tap on its Protection icon.

## Accessing system settings

	Tap on System settings icon on the status bar. / The number of available options depends of firmware version. / When System settings page is displayed it is still possible to check channel’s output state and values (e.g. in this example all outputs are turned off).
	Exit System settings page.