# Firmware upgrade

This chapter describes the procedure for downloading firmware to EEZ BB3 from Linux and Windows operating systems using the USB DFU (Device Firmware Upgrade). Regardless of the PC operating system, the following steps will be required:

	Disconnect any connected loads from all outputs.
	Download the latest firmware available at https://github.com/eez-open/modular-psu-firmware/releases

## Linux
The procedure described below is made on Ubuntu 18.04 LTS and may be somewhat different with other Linux distributions. DFU is an official USB device class specification. It is natively supported by Linux. Therefore, no additional driver will need to be installed.

	Connect BB3 to PC using USB cable. The connector on the BB3 side should be of USB Mini type.
	Turn power off using the power switch on the front panel and keep the BOOT0 switch pressed when power is turned on. The BOOT0 switch is located in the lower left corner of the front panel. If the BOOT0 switch has been pressed long enough, BB3 enters DFU mode and the existing firmware will be deactivated and the Welcome page will not be displayed. The cooling fan has to stop spinning after few seconds.
	You can check if Linux correctly identified BB3, which must be in DFU mode. Open the terminal application and enter: / lsusb / A list of all connected and recognized USB devices will be displayed. If EEZ BB3 is properly connected and enters DFU mode a red marked line should appear. / Bus 004 Device 002: ID 174c:55aa ASMedia Technology Inc. ASM1051E SATA 6Gb/s bridge, ASM1053E SATA 6Gb/s bridge, ASM1153 SATA 3Gb/s bridge / Bus 004 Device 001: ID 1d6b:0003 Linux Foundation 3.0 root hub / Bus 003 Device 002: ID 046d:c077 Logitech, Inc. M105 Optical Mouse / Bus 003 Device 001: ID 1d6b:0002 Linux Foundation 2.0 root hub / Bus 002 Device 001: ID 1d6b:0003 Linux Foundation 3.0 root hub / Bus 001 Device 108: ID 0483:df11 STMicroelectronics STM Device in DFU Mode / Bus 001 Device 103: ID 1a40:0101 Terminus Technology Inc. Hub / Bus 001 Device 003: ID 04d9:1400 Holtek Semiconductor, Inc. PS/2 keyboard + mouse controller / Bus 001 Device 006: ID 8087:0025 Intel Corp. / Bus 001 Device 004: ID 046d:081b Logitech, Inc. Webcam C310 / Bus 001 Device 002: ID 1a40:0101 Terminus Technology Inc. Hub / Bus 001 Device 001: ID 1d6b:0002 Linux Foundation 2.0 root hub
	Install dfu-util (you need sudo access rights): / sudo apt install dfu-util
	Run dfu-util in folder where firmware image is downloaded: / sudo dfu-util -a 0 -D bb3.dfu
	Wait until the firmware image downloading is not completed. A typical dfu-util output is shown below. / dfu-util 0.9 / Copyright 2005-2009 Weston Schmidt, Harald Welte and OpenMoko Inc. / Copyright 2010-2016 Tormod Volden and Stefan Schmidt / This program is Free Software and has ABSOLUTELY NO WARRANTY / Please report bugs to http://sourceforge.net/p/dfu-util/tickets/ / Match vendor ID from file: 0483 / Match product ID from file: 0000 / Opening DFU capable USB device... / ID 0483:df11 / Run-time device DFU version 011a / Claiming USB DFU Interface... / Setting Alternate Setting #0 ... / Determining device status: state = dfuERROR, status = 10 / dfuERROR, clearing status / Determining device status: state = dfuIDLE, status = 0 / dfuIDLE, continuing / DFU mode device DFU version 011a / Device returned transfer size 2048 / DfuSe interface name: "Internal Flash " / file contains 1 DFU images / parsing DFU image 1 / image for alternate setting 0, (2 elements, total size = 901464) / parsing element 1, address = 0x08000000, size = 504 / Download [=========================] 100% 504 bytes / Download done. / parsing element 2, address = 0x08000200, size = 900944 / Download [=========================] 100% 900944 bytes / Download done. / done parsing DfuSe file

## Windows
	Visit the following page / https://www.st.com/en/development-tools/stsw-stm32080.html and select en.stsw-stm32080.zip download. You may require to register and login on the ST web site first.
	Unzip en.stsw-stm32080.zip and start installation of DfuSe_Demo_V3.0.6_Setup.exe
	Connect BB3 to PC using USB cable. The connector on the BB3 side should be of USB Mini type.
	Turn power off using the power switch on the front panel and keep the BOOT0 switch pressed when power is turned on. The BOOT0 switch is located in the lower left corner of the front panel. If the BOOT0 switch has been pressed long enough, BB3 enters DFU mode and the existing firmware will be deactivated and the Welcome page will not be displayed. The cooling fan has to stop spinning after few seconds.
	Check if Windows recognize BB3 as a device in a DFU mode. It should be listed in Device Manager under / Universal Serial Bus controllers section.
	Start the DfuSe Demo, select the Choose button under Upgrade or Verify Action section to load bb3.dfu firmware image. / A message should appear if firmware image is loaded successfully.
	Select the Upgrade button under the same Upgrade or Verify Action section and select Yes when message box with question appears.
	Wait until the firmware upgrade is complete.