# Power modules calibration
This section gives an overview of the calibration features of the power modules. Recommended calibration interval for power modules is 1 year. This will ensure that power modules remains within specification for the next calibration interval.
For optimum calibration results the following condition are recommended:
- The calibration ambient temperature is stable and between 20 °C and 30 °C.
- Ambient relative humidity is less than 80 %
- Allow a one hour warm-up with no load connected
- Use short and thick cables to connect test setups
## Start calibration wizard / The procedure for first calibration and subsequent recalibration is identical. If the module has not yet been calibrated or the calibration data has been deleted (using the SCPI command) the Channel is not calibration message will be displayed. / SCPI / CALibration:CLEar {<password>} / Calibration is a simple process thanks to the calibration wizard that allows you to perform calibration step by step. During the calibration process three set points have to be entered: MIN, MID and MAX. But in practice two point calibration is performed to re-scale the output programmed and measured values by correcting both slope and offset errors. MID point is used only as an additional validation that between entered MIN and MAX values it’s possible to “draw” a line within allowed minor tolerance.

Only one channel can be calibrated at a time. Within the same calibration session both output voltage and current can be calibrated for the currently selected channel. The calibration procedure can be canceled at any step when entered data will be simply ignored. / The calibration process begins by selecting the Calibrate option for the currently selected channel. / The initial calibration password is eezbb3 and should not be confused with a system password that is not set by default. / SCPI / CALibration ON, {<password>}

## Voltage calibration setup / For voltage calibration, disconnect all loads from the power module and connect a digital multimeter (DMM) across the output terminals. / Make sure that the power module is in the CV mode.

### Voltage calibration steps / Voltage MIN point / This is the first step of the calibration wizard. The output voltage on the channel will be set to the default MIN value. This value, as well as those that follow, may change as needed. However, we recommend that the calibration be performed with the values provided. / SCPI / CALibration:VOLTage:LEVel MIN
Set / The numeric keypad will be displayed by selecting this option. The measured value on the external DMM is entered with an arbitrary number of decimal places. / When at least one calibration value is entered, it will be possible to move back and forth between calibration steps. / SCPI / CALibration:VOLTage {<new value>}
Voltage MID point / The second step of the calibration wizard. Repeat the same procedure as above. / SCPI / CALibration:VOLTage:LEVel MID
Voltage MAX point / The third step of the calibration wizard. Repeat the same procedure as above. / Calibration can be completed only after successful completion of all voltage steps. If you want to skip the current calibration, use the right arrow until the remark page is displayed. / SCPI / CALibration:VOLTage:LEVel MAX

## Current calibration setup / For current calibration, connect an appropriate power resistor (less then 5 Ω) in series with DMM set to current measurement. / The Calibration wizard only allows you to calibrate one of the current ranges at a time. All you have to do is skip all the calibration steps that precede the first calibration step of the desired current range.

### Current calibration steps / Power modules such as DCP405 has two current ranges. In this case, the calibration wizard will first show the steps for calibrating the higher range (i.e. 0 – 5 A). / SCPI / CALibration:CURRent:RANGe HIGH / Current [0-5A] MIN point / This is the first step of the high current range calibration. The output current on the channel will be set to the default MIN value.
If power resistor is connected properly the power module will enter CC mode. / SCPI / CALibration:CURrent:LEVel MIN / Current [0-5A] MID point / The second step of the high current range calibration. The measured current on the external DMM should be entered for the set value. / SCPI / CALibration:CURrent:LEVel MID / Current [0-5A] MAX point / The third step of the high current range calibration. Repeat the same procedure as above. / SCPI / CALibration:CURrent:LEVel MAX

After calibrating the high range current, the following steps are for low range. / SCPI / CALibration:CURRent:RANGe LOW / Current [0-50mA] MIN point / This is the first step of the low current range calibration. The output current on the channel will be set to the default MIN value.
Current [0-50mA] MID point / The second step of the low current range calibration. The measured current on the external DMM should be entered for the set value. / SCPI / CALibration:CURrent:LEVel MID / Current [0-50mA] MAX point / The third step of the low current range calibration. Repeat the same procedure as above. / SCPI / CALibration:CURrent:LEVel MAX

Calibration remark / The last parameter that can be set during calibration is remark. The remark has two parts: / a mandatory one that is system defined (a datetime stamp in format yyyymmdd) and / an optional one that can be up to 32 characters long that contains a description of the calibration (eg only voltage is calibrated, etc.)
SCPI / CALibration:REMark {<user remark>}

Save calibration parameters / The final step is the saving of the calibration parameters. When saving is confirmed and all the data entered are within the allowed tolerances, the calibration parameters will be written to the non-volatile memory of the power module; calibration parameters are retained even if the power module is moved to another slot. / SCPI / CALibration:SAVE

Calib. enabled / This option defines whether calibration parameters are used to set and readout voltage and current output values. / It can be enabled when at least one of the calibration ranges (i.e. voltage, current high or low range) is successfully stored. / SCPI / CALibration:STATe ON