# System settings

## System protection settings / Parameters in this section define channel output states on power up and options when a protection mode is tripped to prevent or reduce the possibility of damaging the connected loads. / Additionally, the system password that is used to unlock the display and calibration password can be defined or changed here.
Force disabling of all outputs on power up / Active by default this option ensures that all channel outputs will be switched off on power up to prevent connected loads from being damaged unintentionally. / SCPI / SYSTem:PON:OUTPut:DISable ON
Shutdown when protection tripped / When selected, any protection mode that is tripped on any channel will automatically put the EEZ BB3 in standby mode. / SCPI / SYSTem:POWer:PROTection:TRIP ON
Switch off all outputs when protection tripped / Similar to the previous option except instead of entering standby mode all channel outputs will be turned off. / SCPI / OUTPut:PROTection:COUPle ON
Change password / The system password is used to unlock the display. By default system password is not defined (“”) and it may contain up to 16 characters. Minimum length is 4 characters. The new password is automatically stored in non-volatile memory. / When selected an on-screen keyboard will appear. Characters entered will be replaced with * shortly after entry. The new password has to be entered twice.
SCPI / SYSTem:PASSword:NEW {<old>}, {<new>}
Change calibration password / The calibration password is used to secure access to calibration data. Unlike a system password this password is defined by default (eezbb3) and has to be entered first before new password can be entered. / The calibration password may contain up to 16 characters. Minimum length is 4 characters. / The new password is automatically stored in non-volatile memory. / SCPI / CALibration:PASSword:NEW {<old>}, {<new>}

## Display settings / Display color theme selection, brightness, luminosity, animation duration and status of display can be selected on this page. / An option for entering screen calibration (that is run by default on first start) is also available.
Theme / The EEZ BB3 comes with three color themes that apply to every page. The default theme is Dark.
Anim. duration / Define how long the animation effect lasts when transitioning between pages. Default is 250 ms. Set this parameter to 0 ms if you want to disable the effect.
Brightness / Controls the intensity of the display backlight. The range of the parameter is 1 to 20, where 20 is full intensity and 1 is backlight off. Default value is 10. / SCPI / DISPlay:BRIGhtness {<value>}
Lightness / This parameter defines the lightness of displayed colors if they are represented in so called HSL color space (hue, saturation, lightness). By increasing this value all colors will appear “lighter” or washed away. Similarly, when decreased all colors will appear “darker” or more dull. Lightness is not affected when set to 10. However, due to limitation of the LCD display (color representation is reduced to 16-bit) a default value is set to 5 to effectively compensate color difference. / The effect of changing this parameter can be observed by the color of three screen “widgets” (Background, Status, Pop up).
Turn display off / By turning display off user interaction is no longer possible. To reactivate it, you have to touch and hold anywhere on the display for a short time. / SCPI / DISPlay OFF
Touchscreen calibration / Screen calibration is required to give the firmware the necessary information about its geometry, which will increase the accuracy of data entry. / Initiate this procedure as necessary if you notice that the accuracy of data entry is no longer satisfactory. / SCPI / CALibration:SCReen:INIT
For a successful calibration, you have to use stylus or similar rounded tip object and touch the three dots on the screen as accurately as possible.
After successfully defining the three points, a new page will appear where you can choose between saving calibration data, repeating or canceling the calibration. / On the same screen is possible to test precision by touching the surface beyond the options mentioned.

## Sound settings / Sound / When the beeper is enabled, the EEZ BB3 generates audible sound in any of the following situations: / the power is turned on or off / when error occurs during local (via display) or remote operation / self-test is failed or / any of the protection functions are tripped.
SCPI / SYSTem:BEEPer:STATe {<bool>} / Click sound / Select this option to enable or disable generation of an audible “click” sound as confirmation of selected item on the screen. / SCPI / SYSTem:BEEPer:KEY:STATe {<bool>}

## General trigger settings / The BB3's triggering system allows a change in output state, voltage and current or start internal data logging when receiving a trigger from selected trigger source. / One common triggering system is used to control all available channels, but will only affect those channels that are not in Fixed trigger mode (see Channel trigger settings for channel trigger mode settings).
Source / Bus – enables LAN and serial (via USB) triggering using the *TRG command. When this mode is selected, TRG text will blink when you turn on channels that have triggers defined (e.g. List or Step). / Immediate – the EEZ BB3 executes a complete trigger operation immediately after executing the INITiate command without delay.
Manual – allows you to manually trigger with the User SW (if assigned this function, see User SW) or by selecting the flashing TRG text that will appear on the screen as it shows the example in the picture to the right (Ch2 to Ch4 have triggers defined). / Pin<n> – selects a digital port pin configured as a trigger input. <n> specifies the pin number.

SCPI
TRIGger[:SEQuence]:SOURce {<source>}

Delay
Sets the time delay between the detection of an event on the specified trigger source and the start of any corresponding trigger action on the channel output.

SCPI
TRIGger[:SEQuence]:DELay {<delay>}

Initiate continuously
This option defines whether the trigger system is continuously initiated or not. When set to OFF, the system remains in the IDLE state until it is set to ON or an INITiate:IMMediate command is received.
Once it is set to ON, the trigger system will be initiated and exit the IDLE state. On completion of each trigger cycle, the trigger system immediately commences another trigger cycle without entering the IDLE state.
When this option is set to OFF, the current trigger cycle will be completed before entering the IDLE state. The return to IDLE also occurs as the result of an ABORt or *RST command.

SCPI
INITiate:CONTinuous {<bool>}

## Digital I/O settings / Two protected/buffered digital input pins plus two digital output pins are available on the EEZ BB3 front panel. Their function and polarity can be set on this page which also displays their current state. / Function / Input pins may be assigned with one of the following functions: / Input – The pin is in digital input mode.
Inhibit – When pin is configured as an inhibit input, a true signal at the pin will disable all output channels. If the channel outputs were turned on, all output values will flash in red when entering inhibit mode, as shown in the figure to the right. / Trigger input – When configured as a trigger input, the pin can be selected as the source for trigger signals.

Output pins could be assigned with one of the following functions:
- Output – The pin is in digital output mode.
- Fault – pin functions as an isolated fault output. The fault signal is true when any output is in a protected state (from OCP, OVP, OTP, OPP) or Fan fault is detected.
- Channel ON couple – pin synchronizes channel output state.
- Trigger output – This allows a BUS trigger to be sent to any digital port pin that has been configured as a trigger output. A trigger out pulse is generated when the state is on and a bus trigger is received. A BUS trigger is generated using the *TRG command.
- PWM – (Dout2, pin 4 only) when selected, a square wave will be generated with set frequency from 0.03 Hz to 5 MHz and duty from 0 to 100 %.
SCPI
SYSTem:DIGital:PIN<n>:FUNCtion {<function>}
SYSTem:DIGital:OUTPut:PWM:FREQuency {<pin>}, {<frequency>}
SYSTem:DIGital:OUTPut:PWM:DUTY {<pin>}, {<duty>}

Polarity
- Pos – a logical true signal is a voltage high at the pin. For trigger inputs and outputs, positive means a rising edge.
- Neg – a logical true signal is a voltage low at the pin. For trigger inputs and outputs, negative means a falling edge.

SCPI
SYSTem:DIGital:PIN<n>:POLarity {<polarity>}

State
Current state of all digital inputs and outputs are displayed in this section. If the pin function is not yet defined, the displayed state will be Unassigned.

SCPI
SYSTem:DIGital:INPut:DATA? {<pin>}
SYSTem:DIGital:OUTPut:DATA? {<pin>}

## Date & time / Date / Sets the date of the system clock (RTC). Specify the year, month, and day. / SCPI / SYSTem:DATE {<year>}, {<month>}, {<day>} / Time / Sets the time of the system clock (RTC). Specify the hours, minutes, and seconds.

SCPI
SYSTem:TIME {<hours>}, {<minutes>}, {<seconds>}

Zone
Defines time zone as offset from GMT.

SCPI
SYSTem:TIME:ZONE {<zone>}

DST
Determines Daylight saving time (DST) rules used in your region. The following rules may apply: Europe, USA or Australia.

SCPI
SYSTem:TIME:DST {<rules>}

Format
Sets one of the possible four combination of date, month, year and 12/24 hour that is used for e.g. displaying datetime of the file when accessing it from the File manager.

SCPI
SYSTem:DATE:FORMat {<format>}

Enable / Disable NTP
Internal clock synchronization with NTP server can be enabled or disabled. When enabled, a NTP server entry field will be displayed.
The NTP communication requires active Ethernet connection.
List of most appropriate and available NTP servers is available at ntppool.org

SCPI
SYSTem:COMMunicate:ENABle {<bool>}, NTP

NTP server
NTP service server network address. The EEZ BB3 will try to establish connection with the selected NTP server on every power up (hard reset), when *RST is issued or once per day.

SCPI
SYSTem:COMMunicate:NTP {<server>}

## Encoder & User sw / Confirmation mode / The encoder comes with a tactile switch and if this mode is selected then it will be used to confirm the newly changed value. By definition, the mode is not enabled which means that pressing the switch will instead move the "cursor" to the next editable field.
Acceleration sensitivity moving up/down / The rotation speed of the encoder is measured and defines in what steps the selected value will change. / The reaction (sensitivity) to increasing the rotation speed in one direction or the other (increasing values or decreasing values) can be set by these parameters.

## System temp and fan / Current temp. / The "system" temperature is measured on the AUX-PS module and can be used for over-temperature protection (OTP). / SCPI / MEASure:TEMPerature? AUX

OTP active
Enable or disable over-temperature protection associated with the system (AUX) temperature sensor.

SCPI
SYSTem:TEMPerature:PROTection:STATe

Trip level
Set the over-temperature protection (OTP) value in degrees Celsius (oC). Default value for this sensor is 45 oC.

SCPI
SYSTem:TEMPerature:PROTection {<temperature>}

Trip delay
Specifies how long the temperature should be equal to or higher than the set trip level for the protection to activate.

SCPI
SYSTem:TEMPerature:PROTection {<temperature>}

Clear trip
When protection is activated (tripped), this condition will be latched and further usage of the EEZ BB3 will be disabled until protection is cleared.

SCPI
SYSTem:TEMPerature:PROTection[:HIGH]:CLEar

Fan speed
Current fan speed indicator stated in rpm (revolutions per seconds). The diagnosed fan fault will be also displayed here.

Fan mode
The fan is controlled by default with firmware algorithm (Auto mode) that takes into account measured temperature on all temperature sensors and measured output currents (i.e. the higher the output current, the sooner the fan speed will increase).

Manual mode can be selected for e.g. fan testing purpose when firmware speed control algorithm is completely bypassed.

Set speed to
This option becomes available when manual mode is selected. Default value is 100 % (full speed).

SCPI
SYSTem:FAN:STATus?
SYSTem:FAN:SPEed?

## Ethernet / Enable/disable / Defines activity of the serial interface. When LAN cable is disconnected or network communication is not yet established, Connecting … will be displayed until communication is established again. / SCPI / SYSTem:COMMunicate:ENABle {<bool>}, ETHernet

DHCP
Enable or disable the DHCP mode. In DHCP mode, the DHCP server in the current network assigns network parameters (IP address, DNS address, Gateway address and the Subnet mask). Assigned IP address will we displayed, and cannot be changed.
When DHCP is disabled all network parameters have to be entered manually using the Edit static address as described below.

SCPI
SYSTem:COMMunicate:ETHernet:DHCP {<bool>}

MAC / Defines Ethernet MAC address, a unique identifier assigned to Ethernet interface for communications at the data link layer of a network segment. / Any combination of six hexadecimal values is allowed separated by “-”. The Ethernet connection will work as long as two different machines in the LAN don't have the same MAC address. / SCPI

SYSTem:COMMunicate:ETHernet:MAC {<mac_address>}

SCPI port
Defines Ethernet communication port for SCPI connections. Default port is 5025.

SCPI
SYSTem:COMMunicate:ETHernet:PORT {<number>}

Host name
Defines local area network (LAN) connection unique host name. Host name could contains up to 63 alphanumeric characters. “-” is allowed only if is not used as the first character. Space is not allowed and default name is EEZ-BB3.

SCPI
SYSTem:COMMunicate:ETHernet:PORT {<number>}

### MQTT / When MQTT is enabled, the following parameters should be defined to establish communication: / Host – The name of MQTT server (aka broker) / Port – TCP/IP port for MQTT comminication (standard port is 1883) / Username – login name of a valid account on the MQTT server / Password – login password of a valid account on the MQTT server / Period – MQTT messages publishing frequency

SCPI
SYSTem:COMMunicate:MQTT:SETTings {<address>}, {<port>}, {<user>}, {<password>}, {<period>}

### Static LAN settings (DHCP disabled)
The parameters that follows should be defined when DHCP is not selected. All values should be specified in IPv4 address format (four values from 0 to 255 separated by dots). / IP address / Static local area network (LAN) address of the BB3. / SCPI / SYSTem:COMMunicate:ETHernet:ADDRess {<ip_address>} / DNS / IP address of the DNS (Domain Name Service) server that translates domain names into IP addresses. / SCPI / SYSTem:COMMunicate:ETHernet:DNS {<ip_address>}

Gateway
IP network gateway address for accessing the EEZ BB3 from outside the current sub-network.

SCPI
SYSTem:COMMunicate:ETHernet:GATEway {<ip_address>}

Subnet mask
IP network subnet mask. The subnet mask is used to determine if a client IP address is on the same local subnet.

SCPI
SYSTem:COMMunicate:ETHernet:SMASk {<mask>}

## Serial (via USB) / Enable/disable / Defines activity of the serial (via USB) interface. / When enabled serial communication parameters will be displayed (Baud, Bits, Parity and Stop bits). / SCPI / SYSTem:COMMunicate:ENABle {<bool>}, SERial

## System information / The page summarized various system information and contains a list of all recognized modules. / Total On time / total active time of the EEZ BB3 (i.e. MCU module). Resolution is 1 minute and this information is stored every 10 minutes in non-volatile memory. Therefore it’s possible that up to / 10 minutes is lost after restart caused by a power outage or ungraceful power shutdown.

 SCPI
SYSTem:CPU:INFOrmation:ONTime:TOTal?

Last On time
This query returns time passed after last activation of the EEZ BB3. Resolution is 1 minute and this information is stored every 10 minutes in non-volatile memory. Therefore it’s possible that up to 10 minutes is lost after restart caused by a power outage or system reset.

 SCPI
SYSTem:CPU:INFOrmation:ONTime:LAST?

Temp. AUX
The "system" temperature is measured on the AUX-PS module and can be used for over-temperature protection (OTP). See System temp & fan.

 SCPI
MEASure:TEMPerature? AUX

Fan speed
Current fan speed indicator stated in rpm (revolutions per seconds). Any diagnosed fan faults will be also displayed here. See also System temp & fan.

 SCPI
DIAGnostic:FAN?

Battery
Current voltage of the RTC (Real-time-clock) lithium 3 V coin battery (CR2032, Ø20 x 3.2 mm).

 SCPI
SYSTem:MEASure?

SD card
Information about presence of mass-storage media (i.e. micro SD card).

 SCPI
MMEMory:INFOrmation?

Firmware
The EEZ BB3 (MCU module) firmware version number.

Serial No
The EEZ BB3 (MCU module) serial number.

 SCPI
*IDN?

CPU module
The model name and version of the MCU module.

 SCPI
SYSTem:CPU:MODel?

CH<n> board
The model name and version of all detected modules.

 SCPI
SYSTem:CHANnel:MODel?